## Slide 1: Next Generation Sequencing – An Overview
Olga Vinnere Pettersson, PhD
National Genomics Infrastructure hosted by ScilifeLab,
Uppsala Node (UGC)
Version 5.2.3.d

## Slide 2: Today we will talk about:
History and current state of genomic research
Sequencing technologies:
  Types
  Principles (chemistry & workflow)
  DNA quality requirements
  Choosing the right technology
  Couple of pieces of advise
National Genomics Infrastructure – Sweden
www.robustpm.com

## Slide 3
Massively parallel sequencing (454, Illumina, Life Tech)
Human genome
James Watsons genome
DNA sequencing revolution
Center for Metagenomic
Sequence Analysis (KAW)
Science for Life
Laboratory (SciLifeLab)
Swedish National Infrastructure
for Large-Scale Sequencing
(SNISS)
Notes: Tight schedule, Lunch (wrap) at level 2, Dinner at level 2, speakers hand in USB, Thanks to Anders Andersson

## Slide 5
What is sequencing?

## Slide 6
“In genetics and biochemistry, sequencing means to determine the primary structure (or primary sequence) of an unbranched biopolymer.”
(http://en.wikipedia.org/wiki/Sequencing)
DEFINITION

## Slide 7: Once upon a time…
Fredrik Sanger and Alan Coulson
Chain Termination Sequencing (1977)
Nobel prize 1980
Principle:
SYNTHESIS of DNA is randomly TERMINATED at different points
Separation of fragments that are 1 nucleotide different in size

## Slide 8: Sanger’s sequencing
Lack of OH-group at 3’ position of deoxyribose
!
Fluorescent dye terminators
P32 labelled ddNTPs
Max fragment length – 750 bp

## Slide 9: Maxam & Gilbert Sequencing

## Slide 10: Sequencing genomes using Sanger’s method
Extract & purify genomic DNA
Fragmentation
Make a clone library
Sequence clones
Align sequencies ( -> contigs -> scaffolds)
Close the gaps
Cost/Mb=1000 $, and it takes TIME

## Slide 11: At the very beginning of genome sequencing era…
First genome: virus  X 174 - 5 368 bp (1977)
First organism: Haemophilus influenzae - 1.5 Mb (1995)
First eukaryote: Saccharomyces cerevisiae - 12.4 Mb (1996)
First multicellular organism: Cenorhabditis elegans - 100 MB (1998-2002)
First plant: Arabidopsis thaliana - 157 Mb (2000)

## Slide 12
Just an interesting comparison:
Human genome project, 2007
  Genome of Craig Wenter costs 70 mln $
    Sanger’s sequencing
  Genome of James Watson costs 2 mln $
    454 pyrosequencing
  Ultimate goal: 1000 $ / individual
    Almost there! (1200 $)

## Slide 13: Paradigm change
  From single genes to complete genomes
  From single transcripts to whole transcriptomes
  From single organisms to complex metagenomic pools
  From model organisms to the species you are studying

## Slide 14
IF 2.9
IF 31.6

## Slide 15: Main hazard - DATA ANALYSIS
http://finchtalk.geospiza.com
!
=> More bioinformaticians to people!
$
Sequencing
Data analysis

## Slide 16: Major NGS technologies

## Slide 17: NGS technologies
RIP technologies: Helicos, Polonator, etc.
In development: Tunneling currents, nanopores, etc.
Company | Platform | Amplification | Sequencing method
Roche | 454** | emPCR | Pyrosequencing
Illumina | HiSeq MiSeq | Bridge PCR | Synthesis
LifeTech | SOLiD** | emPCR/ Wildfire | Ligation
LifeTech | Ion Torrent Ion Proton | emPCR | Synthesis (pH)
Pacific Bioscience | RSII | None | Synthesis
Complete genomics | Nanoballs | None | Ligation
Oxford Nanopore* | GridION | None | Flow

## Slide 18: Differences between platforms
Technology: chemistry + signal detection
Run times vary from hours to days
Production range from Mb to Gb
Read length from <100 bp to > 20 Kbp
Accuracy per base from 0.1% to 15%
Cost per base varies

## Slide 19: Roche
Instrument | Yield and run time | Read Length | Error rate | Error type
454 FLX+ | 0.9 GB, 20 hrs | 700 | 1% | Indels
454 FLX Titanium | 0.5 GB, 10 hrs | 450 | 1% | Indels
454 FLX Jr | 0.050 GB, 10 hrs | 400 | 1% | Indels
Main applications:
Microbial genomics and metagenomics
Targeted resequencing

## Slide 20: 454 Titanium GS FLX

## Slide 21: Illumina
Main applications
Whole genome, exome and targeted reseq
Transcriptome analyses
Methylome and ChiPSeq
Rapid targeted resequencing (MiSeq)
Human genome seq (Xten)
Instrument | Yield and run time | Read Length | Error rate | Error type
HiSeq2500 | 120 Gb – 600 Gb 27h or standard run | 100x100 (250x250) | 0.1% | Subst
MiSeq | 540 Mb – 15 Gb (4 – 48 hours) | Up to 350x350 | 0.1% | Subst
HiSeqXten | 800 Gb - 1.8 Tb (3 days) | 150x150 | “ | “

## Slide 22: Illumina

## Slide 23: Illumina reads
5’
3’
Read1
Read2
Index read
5’
3’
Paired-end sequencing
Notes: I sekvensningen utförs 1-4 läsningar beroende på typ av sekvensning.
R1, I1, R2

## Slide 24: Life Technologies SOLiD
Instrument | Yield and run time | Read Length | Error rate | Error type
SOLiD 5500 wildfire | 600 GB, 8 days | 75x35 PE 60x60 MP | 0.01% | A-T Bias
Features
High accuracy due to two-base encoding
True paired-end chemistry - ligation from either end
Mate-pair libraries
Main applications (currently)
ChiPSeq

## Slide 25: SOLiD - ligation

## Slide 26: Life Technologies - Ion Torrent & Ion Proton
Main applications
Microbial and metagenomic sequencing
Targeted resequencing
Clinical sequencing
Chip | Yield - run time | Read Length
PGM 314 | 0.1 GB, 3 hrs | 200 – 400
PGM 316 | 0.5GB, 3 hrs | 200 - 400
PGM 318 | 1 GB, 3 hrs | 200 - 400
P-I | 10 GB | 200

## Slide 27
314 chip 316 chip 318 chip PI chip
10 Mb 100 Mb 1 Gb 10 Gb
virus, bacteria, small eukaryote
eukaryote
200 – 400 bp
200 bp

## Slide 28
316 chip (100 Mbp)
314 chip (10 Mbp)
318 chip (1 Gbp)
IonTorrent Throughput - 400bp
Notes: 28

## Slide 29: Ion Proton - Throughput
We now get 10-16GB data from the PI chip
> 90M reads
~ 150bp read length

## Slide 30: Ion Torrent - H+ ion-sensitive field effect transistors

## Slide 31: Pacific Bioscience
Instrument | Yield and run time | Read Length | Error rate | Error type
RS II | 500 Mb – 1.3 Gb /180 - 240 min SMRTCell | 250 bp – 20 000 bp (70 000 bp) | 15% (on a single passage!) | Insertions, random
Single-Molecule, Real-Time DNA sequencing

## Slide 36
Oxford Nanopore MinION
Reads up to 100k
1D and 2D reads
15-40% error rate
Life time 5 days

## Slide 37
 | Illumina HiSeq | Illumina MiSeq | SOLiD Wildfire | Ion Torrent | Ion Proton | PacBio
Read length | 100 + 100 bp (150+150 bp) | 250 + 250 bp (350+350 bp) | 75 bp | 200 bp 400 bp (500 bp) | 150 bp 200 bp | 250 bp – 40 Kbp
WGS: - human - small | ++++ ++ | +++ | (+) (+) | +++ | + ++ | (+++++) +++++
De novo | +++ | ++ |  | +++ | ++ | +++++
RNA-seq miRNA | +++ +++ |  | +++ +++ |  | +++ +++ | +++*
ChIP | +++ |  | ++++ |  |  | 
Amplicon** | +++ | +++ |  | +++ | +++ | ++++
Metylation | +++ |  |  |  |  | ++++***
Target re-seq | ++ | +++ | (+) |  | +++ | +++
Exome | +++ |  | (+) |  | ++++ | (+)
* Only whole-length transcripts; ** Depends on amplicon size; *** Only prokaryotes

## Slide 38: What is The BEST?

## Slide 39: Garbage in – garbage out: sequencing success to 90% depends on the sample quality

## Slide 40: Making a NGS library
DNA QC – paramount importance
Sharing & size selection
Ligation of sequencing adaptors, technology specific
Amplification

## Slide 41: What do absorption ratios tell us?
Pure DNA 260/280: 1.8 – 2.0
< 1.8:
Too little DNA compared to other components of the solution; presence of organic contaminants: proteins and phenol; glycogen - absorb at 280 nm.
> 2.0:
High share of RNA.
Pure DNA 260/230: 2.0 – 2.2
<2.0:
Salt contamination, humic acids, peptides, aromatic compounds, polyphenols, urea, guanidine, thiocyanates (latter three are common kit components) – absorb at 230 nm.
>2.2:
High share of RNA, very high share of phenol, high turbidity, dirty instrument.
Photometrically active contaminants:
phenol, polyphenols, EDTA, thiocyanate, protein,
RNA, nucleotides (fragments below 5 bp)

## Slide 42: Sample prep: general recommendations
Treat DNA as a crystal vase: it is fragile when in solution
As soon as DNA is released from the cells – use wide-bore tips
Limit pipetting to minimum
Never vortex!
Do not heat above 65°C
Reduce amount of freeze-thaw cycles to minimum
Make several aliquotes of the stock DNA

## Slide 43: How to make a correct measurement
Thaw DNA completely
Mix gently (never vortex!)
Put the sample on a thermoblock: 37°C, 15-30 min
Mix gently
Dilute 1:100 (if HMW)
Mix gently
Make a measurement with an appropriate blank
Low concentration
High concentration
DNA
solution

## Slide 44: Post-extraction cleanup
Protein contamination
- Apply phenol-chloroform
Phenol carry-over or
overloaded sample?
RNA contamination
- Apply RNase, followed
by phenol-chloroform
extraction
If unsure, make dilution series.
If problem persists – try MoBio clean-up kit,
or re-extract DNA

## Slide 45: Input QC control at NGI:
Qubit for DNA
  Measures content of dsDNA only
  Nanodrop & NanoVue overestimate concentrations up to 300%!
Bioanalyzer for RNA and amplicons
  RNA: RIN values and concentrations
  Amplicons: size distribution (extremely important!)

## Slide 46
Example 2: several sizes,
fractionation is needed
=> we HAVE to make several libraries
Example 3: broad peak;
size selection is needed
Bioanalyzer: amplicon size check
FOR ANY NGS TECHNOLOGY
Size difference among fragments must not
exceed 80 bp (optimally 50 bp)
Reason – preferential amplification of short fragments
Example 1: OK size distribution

## Slide 47: Sample prep: take home message
PCR-quality sample and
NGS-quality sample
are two completely different things

## Slide 48: NGS technologies - SUMMARY
Platform | Read length | Accuracy | Projects / applications
454 | Medium | Homo- polymer runs | Microbial + targeted reseq
HiSeq MiSeq | Short Medium | High | Whole genome + transcriptome seq, exome
SOLiD | Short | High | Whole genome + transcriptome seq, exome
Ion Torrent | Medium | High | Microbial + targeted reseq
Ion Proton | Short/Medium | High | Exome, transcriptome, genome
PacBio | Long | Low – ultra high* | Microbial + targeted reseq Gap closure & scaffolding
MinION | Long | Low | Gap closure, scaffolding structural variants

## Slide 49
Check list:
Have others done similar work?
Is your methodology sound? Sample size? Repetitions?
Is there people to analyze the data?
Is there computer capacity to analyze the data?
Will you be able to publish NGS data by yourself?
PLEASE consult the sequencing facility & BILS
PRIOR to onset of your project!

## Slide 50
Common pitfalls and a piece of advise:
If you give us low quality DNA/RNA - expect low quality data
If you give us too little DNA/RNA – expect biased data
Do not try to do everything by yourself
Make sure there is a dedicated bioinformatician available
Never underestimate time and money needed for data analysis
Google often!
Use online forums, e.g. SeqAnswers.com

## Slide 51: Summary
Progress is FAST- keep yourselves updated!
Chose technology based on:
  What is most feasible
  What is most accessible
  What is most cost-effective
  SciLifeLab Genomics & Bioinformatics are here for you!

## Slide 54: National Genomics Infrastructure
Mid 2010
SciLifeLab, Stockholm
SciLifeLab, Uppsala
Uppmax, Uppsala

## Slide 55
Projects at CMS
3. Access to genomics platform
Portal project flow
NGI Project coordinators meet every second day via Skype
Project is then assigned to a certain node and a coordinator contacts the PI
Project distribution is based on:
Wish of PI
Type of sequencing technology
Type of application
Queue at technology platforms
Ulrika Liljedahl
SNP&SEQ
Uppsala node
Mattias Ormestad
Stockholm Node
Olga Vinnere Pettersson
UGC
Uppsala Node

## Slide 56
Illumina HiSeq 2000/2500 17
Illumina MiSeq 3
Illumina X 10
Life Technologies SOLiD 5500wildfire 1
Life Technologies Ion Torrent 2
Life Technologies Ion Proton 6
Life Technologies Sanger ABI3730 2
Pacific Biosciences RSII 2
Argus Whole Genome Mapping System 1
NGI Equipment
One of 5 best-equipped NGS sites in Europe

## Slide 57
Projects at CMS
3. Access to genomics platform
Project meeting
What we can help you with:
Design your experiment based on the scientific question.
Chose the best suited application for your project.
Find the most optimal sequencing setup.
Answer all questions about our technologies and applications, as well as bioinformatics.
In special cases, we can give extra-support with bioinformatics analysis – development of novel methods and applications

## Slide 58
Downstream Data Analysis
Bioinformatics competence IS present in research group
Bioinformatics competence IS NOT present in research group
BILS:
Bioinformatics Infrastructure
for Life
Sciences
WABI:
Wallenberg
Advanced
Bioinformatics
Initiative
Short-term commitment
Long-term commitment
Cooperation with platform personnel: R&D
Co-authorship